## run_notes
Plan for Tuesday 2025-06-09
Go to UNM, set out bucket of water to equilibrate to room temp
Pack at UNM
four clean water jugs
two white buckets
pump + tubing
sharpie
parafilm
rubber bands
field bag
gloves?
Go to sites and collect water as normal
Take it back to campus
Put jugs in buckets of room temp water (for how long?)
Make remaining calibration standard and five QCs
Make matrix spikes of each well (not filtered)
10 mg/L (times 4)
5 mg/L (times 4)
2 mg/L (times 4)
1 mg/L (times 4)
0.5 mg/L (times 4)
Put in buckets of room temp water to equilibrate (for how long?)
Procedure for NPOC analysis:
Rinse the filter with the standard amount (20 mL?)
Filter ~ 20 mL of sample into the vial
Not acidifying!
Procedure for filter rinse test:
CHECK IF THERE ARE ENOUGH FILTER HOLDERS
Have nanopure in a beaker
Draw up ~ 50 mL of sample with a clean syringe
Attach filter holder + filter
Push through 5 mL of sample into vial 1, then 5 mL of sample into vial 2, etc.
Repeat the above step with four filter holders + filters for a total of 20 mL per vial
Repeat the above four steps with another syringe-full of 50 mL of nanopure, reusing the same filters

## raw_data
[Header]
System | TOC-L
S/W Version | 1
User
Date/Time | 2025-06-11 12:47
[System]
Instrument Options | TOC/ASI/IC Unit/TN/
[Data]
Type | Anal. | Sample Name | Sample ID | Date / Time | Spl. No. | Inj. No. | Analysis(Inj.) | Area | Mean Area | Conc. | SD Area | CV Area | Vial | Excluded | Inj. Vol. | Auto. Dil.
Unknown | NPOC | 0position | Untitled | 2025-06-10 18:08 | 1 | 1 | NPOC | 0.8239 | 0.8424 | 0 | 0.02616 | 3.11 | 0 | 0 | 100 | 1
Unknown | NPOC | 0position | Untitled | 2025-06-10 18:11 | 1 | 2 | NPOC | 0.8609 | 0.8424 | 0 | 0.02616 | 3.11 | 0 | 0 | 100 | 1
Unknown | NPOC | nanopure | Untitled | 2025-06-10 18:20 | 1 | 1 | NPOC | 0.4395 | 0.4632 | 0 | 0.03345 | 7.22 | 1 | 0 | 100 | 1
Unknown | NPOC | nanopure | Untitled | 2025-06-10 18:22 | 1 | 2 | NPOC | 0.4868 | 0.4632 | 0 | 0.03345 | 7.22 | 1 | 0 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 18:34 | 1 | 1 | NPOC | 138 | 137.3 | 20 | 0.9899 | 0.72 | 2 | 0 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 18:38 | 1 | 2 | NPOC | 136.6 | 137.3 | 20 | 0.9899 | 0.72 | 2 | 0 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 18:49 | 2 | 1 | NPOC | 69.53 | 69.06 | 10 | 0.6576 | 0.95 | 3 | 0 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 18:52 | 2 | 2 | NPOC | 68.6 | 69.06 | 10 | 0.6576 | 0.95 | 3 | 0 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 19:03 | 3 | 1 | NPOC | 35.56 | 34.2 | 5 | 0.2051 | 0.6 | 4 | 1 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 19:06 | 3 | 2 | NPOC | 34.35 | 34.2 | 5 | 0.2051 | 0.6 | 4 | 0 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 19:09 | 3 | 3 | NPOC | 34.06 | 34.2 | 5 | 0.2051 | 0.6 | 4 | 0 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 19:19 | 4 | 1 | NPOC | 14.35 | 14.19 | 2 | 0.2333 | 1.65 | 5 | 0 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 19:21 | 4 | 2 | NPOC | 14.02 | 14.19 | 2 | 0.2333 | 1.65 | 5 | 0 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 19:31 | 5 | 1 | NPOC | 7.531 | 7.179 | 1 | 0.04384 | 0.61 | 6 | 1 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 19:34 | 5 | 2 | NPOC | 7.148 | 7.179 | 1 | 0.04384 | 0.61 | 6 | 0 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 19:36 | 5 | 3 | NPOC | 7.21 | 7.179 | 1 | 0.04384 | 0.61 | 6 | 0 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 19:46 | 6 | 1 | NPOC | 3.976 | 3.66 | 0.5 | 0.111 | 3.03 | 7 | 1 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 19:48 | 6 | 2 | NPOC | 3.581 | 3.66 | 0.5 | 0.111 | 3.03 | 7 | 0 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 19:50 | 6 | 3 | NPOC | 3.738 | 3.66 | 0.5 | 0.111 | 3.03 | 7 | 0 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 20:00 | 7 | 1 | NPOC | 0.2326 | 0.3301 | 0 | 0.1379 | 41.77 | 8 | 0 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 20:02 | 7 | 2 | NPOC | 0 | 0.3301 | 0 | 0.1379 | 41.77 | 8 | 1 | 100 | 1
Standard | NPOC | NPOC_CAL | NPOC_CAL | 2025-06-10 20:04 | 7 | 3 | NPOC | 0.4276 | 0.3301 | 0 | 0.1379 | 41.77 | 8 | 0 | 100 | 1
Unknown | NPOC | Vial_19 | Untitled | 2025-06-10 20:16 | 1 | 1 | NPOC | 91.04 | 88.43 | 0 | 0.4667 | 0.53 | 19 | 1 | 100 | 1
Unknown | NPOC | Vial_19 | Untitled | 2025-06-10 20:19 | 1 | 2 | NPOC | 88.1 | 88.43 | 0 | 0.4667 | 0.53 | 19 | 0 | 100 | 1
Unknown | NPOC | Vial_19 | Untitled | 2025-06-10 20:22 | 1 | 3 | NPOC | 88.76 | 88.43 | 0 | 0.4667 | 0.53 | 19 | 0 | 100 | 1
Unknown | NPOC | Vial_20 | Untitled | 2025-06-10 20:33 | 1 | 1 | NPOC | 59.34 | 56.79 | 0 | 0.3677 | 0.65 | 20 | 1 | 100 | 1
Unknown | NPOC | Vial_20 | Untitled | 2025-06-10 20:36 | 1 | 2 | NPOC | 57.05 | 56.79 | 0 | 0.3677 | 0.65 | 20 | 0 | 100 | 1
Unknown | NPOC | Vial_20 | Untitled | 2025-06-10 20:40 | 1 | 3 | NPOC | 56.53 | 56.79 | 0 | 0.3677 | 0.65 | 20 | 0 | 100 | 1
Unknown | NPOC | Vial_21 | Untitled | 2025-06-10 20:51 | 1 | 1 | NPOC | 40.6 | 38.2 | 0 | 0.1556 | 0.41 | 21 | 1 | 100 | 1
Unknown | NPOC | Vial_21 | Untitled | 2025-06-10 20:54 | 1 | 2 | NPOC | 38.31 | 38.2 | 0 | 0.1556 | 0.41 | 21 | 0 | 100 | 1
Unknown | NPOC | Vial_21 | Untitled | 2025-06-10 20:57 | 1 | 3 | NPOC | 38.09 | 38.2 | 0 | 0.1556 | 0.41 | 21 | 0 | 100 | 1
Unknown | NPOC | Vial_22 | Untitled | 2025-06-10 21:08 | 1 | 1 | NPOC | 32.97 | 30.55 | 0 | 0.8132 | 2.66 | 22 | 1 | 100 | 1
Unknown | NPOC | Vial_22 | Untitled | 2025-06-10 21:11 | 1 | 2 | NPOC | 31.12 | 30.55 | 0 | 0.8132 | 2.66 | 22 | 0 | 100 | 1
Unknown | NPOC | Vial_22 | Untitled | 2025-06-10 21:14 | 1 | 3 | NPOC | 29.97 | 30.55 | 0 | 0.8132 | 2.66 | 22 | 0 | 100 | 1
Unknown | NPOC | Vial_23 | Untitled | 2025-06-10 21:24 | 1 | 1 | NPOC | 30.71 | 28.28 | 0 | 0.5869 | 2.08 | 23 | 1 | 100 | 1
Unknown | NPOC | Vial_23 | Untitled | 2025-06-10 21:27 | 1 | 2 | NPOC | 28.69 | 28.28 | 0 | 0.5869 | 2.08 | 23 | 0 | 100 | 1
Unknown | NPOC | Vial_23 | Untitled | 2025-06-10 21:30 | 1 | 3 | NPOC | 27.86 | 28.28 | 0 | 0.5869 | 2.08 | 23 | 0 | 100 | 1
Unknown | NPOC | Vial_24 | Untitled | 2025-06-10 21:41 | 1 | 1 | NPOC | 26.73 | 23.82 | 0 | 0.1131 | 0.47 | 24 | 1 | 100 | 1
Unknown | NPOC | Vial_24 | Untitled | 2025-06-10 21:44 | 1 | 2 | NPOC | 23.9 | 23.82 | 0 | 0.1131 | 0.47 | 24 | 0 | 100 | 1
Unknown | NPOC | Vial_24 | Untitled | 2025-06-10 21:47 | 1 | 3 | NPOC | 23.74 | 23.82 | 0 | 0.1131 | 0.47 | 24 | 0 | 100 | 1
Unknown | NPOC | Vial_9_QC1 | Untitled | 2025-06-10 21:58 | 1 | 1 | NPOC | 72.79 | 69.55 | 0 | 0.2333 | 0.34 | 9 | 1 | 100 | 1
Unknown | NPOC | Vial_9_QC1 | Untitled | 2025-06-10 22:01 | 1 | 2 | NPOC | 69.71 | 69.55 | 0 | 0.2333 | 0.34 | 9 | 0 | 100 | 1
Unknown | NPOC | Vial_9_QC1 | Untitled | 2025-06-10 22:04 | 1 | 3 | NPOC | 69.38 | 69.55 | 0 | 0.2333 | 0.34 | 9 | 0 | 100 | 1
Unknown | NPOC | Vial_14_blank1 | Untitled | 2025-06-10 22:14 | 1 | 1 | NPOC | 0.6249 | 0.5861 | 0 | 0.0548 | 9.35 | 14 | 0 | 100 | 1
Unknown | NPOC | Vial_14_blank1 | Untitled | 2025-06-10 22:16 | 1 | 2 | NPOC | 0.5474 | 0.5861 | 0 | 0.0548 | 9.35 | 14 | 0 | 100 | 1
Unknown | NPOC | Vial_43_SLOC_0_REP | Untitled | 2025-06-10 22:27 | 1 | 1 | NPOC | 22.96 | 22.76 | 0 | 0.2899 | 1.27 | 43 | 0 | 100 | 1
Unknown | NPOC | Vial_43_SLOC_0_REP | Untitled | 2025-06-10 22:30 | 1 | 2 | NPOC | 22.55 | 22.76 | 0 | 0.2899 | 1.27 | 43 | 0 | 100 | 1
Unknown | NPOC | Vial_44_SLOC_10_REP | Untitled | 2025-06-10 22:42 | 1 | 1 | NPOC | 94.55 | 91.35 | 0 | 0.1131 | 0.12 | 44 | 1 | 100 | 1
Unknown | NPOC | Vial_44_SLOC_10_REP | Untitled | 2025-06-10 22:45 | 1 | 2 | NPOC | 91.43 | 91.35 | 0 | 0.1131 | 0.12 | 44 | 0 | 100 | 1
Unknown | NPOC | Vial_44_SLOC_10_REP | Untitled | 2025-06-10 22:49 | 1 | 3 | NPOC | 91.27 | 91.35 | 0 | 0.1131 | 0.12 | 44 | 0 | 100 | 1
Unknown | NPOC | Vial_25 | Untitled | 2025-06-10 23:00 | 1 | 1 | NPOC | 86.13 | 82 | 0 | 0.4313 | 0.53 | 25 | 1 | 100 | 1
Unknown | NPOC | Vial_25 | Untitled | 2025-06-10 23:04 | 1 | 2 | NPOC | 81.7 | 82 | 0 | 0.4313 | 0.53 | 25 | 0 | 100 | 1
Unknown | NPOC | Vial_25 | Untitled | 2025-06-10 23:07 | 1 | 3 | NPOC | 82.31 | 82 | 0 | 0.4313 | 0.53 | 25 | 0 | 100 | 1
Unknown | NPOC | Vial_26 | Untitled | 2025-06-10 23:18 | 1 | 1 | NPOC | 51.25 | 47.88 | 0 | 0.3323 | 0.69 | 26 | 1 | 100 | 1
Unknown | NPOC | Vial_26 | Untitled | 2025-06-10 23:22 | 1 | 2 | NPOC | 48.11 | 47.88 | 0 | 0.3323 | 0.69 | 26 | 0 | 100 | 1
Unknown | NPOC | Vial_26 | Untitled | 2025-06-10 23:25 | 1 | 3 | NPOC | 47.64 | 47.88 | 0 | 0.3323 | 0.69 | 26 | 0 | 100 | 1
Unknown | NPOC | Vial_27 | Untitled | 2025-06-10 23:36 | 1 | 1 | NPOC | 29.91 | 28.41 | 0 | 0.1414 | 0.5 | 27 | 1 | 100 | 1
Unknown | NPOC | Vial_27 | Untitled | 2025-06-10 23:38 | 1 | 2 | NPOC | 28.31 | 28.41 | 0 | 0.1414 | 0.5 | 27 | 0 | 100 | 1
Unknown | NPOC | Vial_27 | Untitled | 2025-06-10 23:42 | 1 | 3 | NPOC | 28.51 | 28.41 | 0 | 0.1414 | 0.5 | 27 | 0 | 100 | 1
Unknown | NPOC | Vial_28 | Untitled | 2025-06-10 23:52 | 1 | 1 | NPOC | 22.81 | 20.77 | 0 | 0.03536 | 0.17 | 28 | 1 | 100 | 1
Unknown | NPOC | Vial_28 | Untitled | 2025-06-10 23:55 | 1 | 2 | NPOC | 20.79 | 20.77 | 0 | 0.03536 | 0.17 | 28 | 0 | 100 | 1
Unknown | NPOC | Vial_28 | Untitled | 2025-06-10 23:58 | 1 | 3 | NPOC | 20.74 | 20.77 | 0 | 0.03536 | 0.17 | 28 | 0 | 100 | 1
Unknown | NPOC | Vial_29 | Untitled | 2025-06-11 00:08 | 1 | 1 | NPOC | 18.72 | 16.89 | 0 | 0.4667 | 2.76 | 29 | 1 | 100 | 1
Unknown | NPOC | Vial_29 | Untitled | 2025-06-11 00:11 | 1 | 2 | NPOC | 17.22 | 16.89 | 0 | 0.4667 | 2.76 | 29 | 0 | 100 | 1
Unknown | NPOC | Vial_29 | Untitled | 2025-06-11 00:14 | 1 | 3 | NPOC | 16.56 | 16.89 | 0 | 0.4667 | 2.76 | 29 | 0 | 100 | 1
Unknown | NPOC | Vial_30 | Untitled | 2025-06-11 00:24 | 1 | 1 | NPOC | 15.46 | 13.77 | 0 | 0.02828 | 0.21 | 30 | 1 | 100 | 1
Unknown | NPOC | Vial_30 | Untitled | 2025-06-11 00:27 | 1 | 2 | NPOC | 13.75 | 13.77 | 0 | 0.02828 | 0.21 | 30 | 0 | 100 | 1
Unknown | NPOC | Vial_30 | Untitled | 2025-06-11 00:30 | 1 | 3 | NPOC | 13.79 | 13.77 | 0 | 0.02828 | 0.21 | 30 | 0 | 100 | 1
Unknown | NPOC | Vial_10_QC2 | Untitled | 2025-06-11 00:41 | 1 | 1 | NPOC | 74.54 | 71.44 | 0 | 0.09192 | 0.13 | 10 | 1 | 100 | 1
Unknown | NPOC | Vial_10_QC2 | Untitled | 2025-06-11 00:44 | 1 | 2 | NPOC | 71.37 | 71.44 | 0 | 0.09192 | 0.13 | 10 | 0 | 100 | 1
Unknown | NPOC | Vial_10_QC2 | Untitled | 2025-06-11 00:47 | 1 | 3 | NPOC | 71.5 | 71.44 | 0 | 0.09192 | 0.13 | 10 | 0 | 100 | 1
Unknown | NPOC | Vial_15_blank2 | Untitled | 2025-06-11 00:57 | 1 | 1 | NPOC | 0.7162 | 0.4273 | 0 | 0.06116 | 14.32 | 15 | 1 | 100 | 1
Unknown | NPOC | Vial_15_blank2 | Untitled | 2025-06-11 00:59 | 1 | 2 | NPOC | 0.4705 | 0.4273 | 0 | 0.06116 | 14.32 | 15 | 0 | 100 | 1
Unknown | NPOC | Vial_15_blank2 | Untitled | 2025-06-11 01:01 | 1 | 3 | NPOC | 0.384 | 0.4273 | 0 | 0.06116 | 14.32 | 15 | 0 | 100 | 1
Unknown | NPOC | Vial_45_SLOW_0_REP | Untitled | 2025-06-11 01:12 | 1 | 1 | NPOC | 12.83 | 12.7 | 0 | 0.1838 | 1.45 | 45 | 0 | 100 | 1
Unknown | NPOC | Vial_45_SLOW_0_REP | Untitled | 2025-06-11 01:14 | 1 | 2 | NPOC | 12.57 | 12.7 | 0 | 0.1838 | 1.45 | 45 | 0 | 100 | 1
Unknown | NPOC | Vial_46_SLOW_10_REP | Untitled | 2025-06-11 01:26 | 1 | 1 | NPOC | 83.89 | 80.74 | 0 | 0.502 | 0.62 | 46 | 1 | 100 | 1
Unknown | NPOC | Vial_46_SLOW_10_REP | Untitled | 2025-06-11 01:30 | 1 | 2 | NPOC | 81.09 | 80.74 | 0 | 0.502 | 0.62 | 46 | 0 | 100 | 1
Unknown | NPOC | Vial_46_SLOW_10_REP | Untitled | 2025-06-11 01:33 | 1 | 3 | NPOC | 80.38 | 80.74 | 0 | 0.502 | 0.62 | 46 | 0 | 100 | 1
Unknown | NPOC | Vial_31 | Untitled | 2025-06-11 01:44 | 1 | 1 | NPOC | 102.9 | 101.5 | 0 | 1.98 | 1.95 | 31 | 0 | 100 | 1
Unknown | NPOC | Vial_31 | Untitled | 2025-06-11 01:47 | 1 | 2 | NPOC | 100.1 | 101.5 | 0 | 1.98 | 1.95 | 31 | 0 | 100 | 1
Unknown | NPOC | Vial_32 | Untitled | 2025-06-11 01:58 | 1 | 1 | NPOC | 70.71 | 66.84 | 0 | 1.054 | 1.58 | 32 | 1 | 100 | 1
Unknown | NPOC | Vial_32 | Untitled | 2025-06-11 02:01 | 1 | 2 | NPOC | 67.58 | 66.84 | 0 | 1.054 | 1.58 | 32 | 0 | 100 | 1
Unknown | NPOC | Vial_32 | Untitled | 2025-06-11 02:05 | 1 | 3 | NPOC | 66.09 | 66.84 | 0 | 1.054 | 1.58 | 32 | 0 | 100 | 1
Unknown | NPOC | Vial_33 | Untitled | 2025-06-11 02:15 | 1 | 1 | NPOC | 51.48 | 46.77 | 0 | 0.792 | 1.69 | 33 | 1 | 100 | 1
Unknown | NPOC | Vial_33 | Untitled | 2025-06-11 02:19 | 1 | 2 | NPOC | 47.33 | 46.77 | 0 | 0.792 | 1.69 | 33 | 0 | 100 | 1
Unknown | NPOC | Vial_33 | Untitled | 2025-06-11 02:22 | 1 | 3 | NPOC | 46.21 | 46.77 | 0 | 0.792 | 1.69 | 33 | 0 | 100 | 1
Unknown | NPOC | Vial_34 | Untitled | 2025-06-11 02:32 | 1 | 1 | NPOC | 42.87 | 40.36 | 0 | 0.0495 | 0.12 | 34 | 1 | 100 | 1
Unknown | NPOC | Vial_34 | Untitled | 2025-06-11 02:35 | 1 | 2 | NPOC | 40.39 | 40.36 | 0 | 0.0495 | 0.12 | 34 | 0 | 100 | 1
Unknown | NPOC | Vial_34 | Untitled | 2025-06-11 02:38 | 1 | 3 | NPOC | 40.32 | 40.36 | 0 | 0.0495 | 0.12 | 34 | 0 | 100 | 1
Unknown | NPOC | Vial_35 | Untitled | 2025-06-11 02:49 | 1 | 1 | NPOC | 39.53 | 36.56 | 0 | 0.3818 | 1.04 | 35 | 1 | 100 | 1
Unknown | NPOC | Vial_35 | Untitled | 2025-06-11 02:52 | 1 | 2 | NPOC | 36.83 | 36.56 | 0 | 0.3818 | 1.04 | 35 | 0 | 100 | 1
Unknown | NPOC | Vial_35 | Untitled | 2025-06-11 02:55 | 1 | 3 | NPOC | 36.29 | 36.56 | 0 | 0.3818 | 1.04 | 35 | 0 | 100 | 1
Unknown | NPOC | Vial_36 | Untitled | 2025-06-11 03:06 | 1 | 1 | NPOC | 37.71 | 33.78 | 0 | 0.9192 | 2.72 | 36 | 1 | 100 | 1
Unknown | NPOC | Vial_36 | Untitled | 2025-06-11 03:09 | 1 | 2 | NPOC | 34.43 | 33.78 | 0 | 0.9192 | 2.72 | 36 | 0 | 100 | 1
Unknown | NPOC | Vial_36 | Untitled | 2025-06-11 03:12 | 1 | 3 | NPOC | 33.13 | 33.78 | 0 | 0.9192 | 2.72 | 36 | 0 | 100 | 1
Unknown | NPOC | Vial_11_QC3 | Untitled | 2025-06-11 03:24 | 1 | 1 | NPOC | 75.36 | 71.33 | 0 | 0.5233 | 0.73 | 11 | 1 | 100 | 1
Unknown | NPOC | Vial_11_QC3 | Untitled | 2025-06-11 03:27 | 1 | 2 | NPOC | 71.7 | 71.33 | 0 | 0.5233 | 0.73 | 11 | 0 | 100 | 1
Unknown | NPOC | Vial_11_QC3 | Untitled | 2025-06-11 03:30 | 1 | 3 | NPOC | 70.96 | 71.33 | 0 | 0.5233 | 0.73 | 11 | 0 | 100 | 1
Unknown | NPOC | Vial_16_blank3 | Untitled | 2025-06-11 03:40 | 1 | 1 | NPOC | 0.9864 | 0.6499 | 0 | 0.09129 | 14.05 | 16 | 1 | 100 | 1
Unknown | NPOC | Vial_16_blank3 | Untitled | 2025-06-11 03:42 | 1 | 2 | NPOC | 0.7144 | 0.6499 | 0 | 0.09129 | 14.05 | 16 | 0 | 100 | 1
Unknown | NPOC | Vial_16_blank3 | Untitled | 2025-06-11 03:44 | 1 | 3 | NPOC | 0.5853 | 0.6499 | 0 | 0.09129 | 14.05 | 16 | 0 | 100 | 1
Unknown | NPOC | Vial_47_VDOS_0_REP | Untitled | 2025-06-11 03:55 | 1 | 1 | NPOC | 31.68 | 31.76 | 0 | 0.1061 | 0.33 | 47 | 0 | 100 | 1
Unknown | NPOC | Vial_47_VDOS_0_REP | Untitled | 2025-06-11 03:58 | 1 | 2 | NPOC | 31.83 | 31.76 | 0 | 0.1061 | 0.33 | 47 | 0 | 100 | 1
Unknown | NPOC | Vial_48_VDOS_10_REP | Untitled | 2025-06-11 04:09 | 1 | 1 | NPOC | 103 | 100.4 | 0 | 0.7707 | 0.77 | 48 | 1 | 100 | 1
Unknown | NPOC | Vial_48_VDOS_10_REP | Untitled | 2025-06-11 04:12 | 1 | 2 | NPOC | 99.81 | 100.4 | 0 | 0.7707 | 0.77 | 48 | 0 | 100 | 1
Unknown | NPOC | Vial_48_VDOS_10_REP | Untitled | 2025-06-11 04:16 | 1 | 3 | NPOC | 100.9 | 100.4 | 0 | 0.7707 | 0.77 | 48 | 0 | 100 | 1
Unknown | NPOC | Vial_37 | Untitled | 2025-06-11 04:27 | 1 | 1 | NPOC | 99.18 | 95.62 | 0 | 0.8415 | 0.88 | 37 | 1 | 100 | 1
Unknown | NPOC | Vial_37 | Untitled | 2025-06-11 04:31 | 1 | 2 | NPOC | 96.21 | 95.62 | 0 | 0.8415 | 0.88 | 37 | 0 | 100 | 1
Unknown | NPOC | Vial_37 | Untitled | 2025-06-11 04:34 | 1 | 3 | NPOC | 95.02 | 95.62 | 0 | 0.8415 | 0.88 | 37 | 0 | 100 | 1
Unknown | NPOC | Vial_38 | Untitled | 2025-06-11 04:45 | 1 | 1 | NPOC | 67.14 | 63.85 | 0 | 0.4808 | 0.75 | 38 | 1 | 100 | 1
Unknown | NPOC | Vial_38 | Untitled | 2025-06-11 04:49 | 1 | 2 | NPOC | 63.51 | 63.85 | 0 | 0.4808 | 0.75 | 38 | 0 | 100 | 1
Unknown | NPOC | Vial_38 | Untitled | 2025-06-11 04:52 | 1 | 3 | NPOC | 64.19 | 63.85 | 0 | 0.4808 | 0.75 | 38 | 0 | 100 | 1
Unknown | NPOC | Vial_39 | Untitled | 2025-06-11 05:03 | 1 | 1 | NPOC | 47.08 | 44.45 | 0 | 0.1697 | 0.38 | 39 | 1 | 100 | 1
Unknown | NPOC | Vial_39 | Untitled | 2025-06-11 05:06 | 1 | 2 | NPOC | 44.33 | 44.45 | 0 | 0.1697 | 0.38 | 39 | 0 | 100 | 1
Unknown | NPOC | Vial_39 | Untitled | 2025-06-11 05:09 | 1 | 3 | NPOC | 44.57 | 44.45 | 0 | 0.1697 | 0.38 | 39 | 0 | 100 | 1
Unknown | NPOC | Vial_40 | Untitled | 2025-06-11 05:20 | 1 | 1 | NPOC | 41.06 | 38.48 | 0 | 0.3748 | 0.97 | 40 | 1 | 100 | 1
Unknown | NPOC | Vial_40 | Untitled | 2025-06-11 05:23 | 1 | 2 | NPOC | 38.75 | 38.48 | 0 | 0.3748 | 0.97 | 40 | 0 | 100 | 1
Unknown | NPOC | Vial_40 | Untitled | 2025-06-11 05:26 | 1 | 3 | NPOC | 38.22 | 38.48 | 0 | 0.3748 | 0.97 | 40 | 0 | 100 | 1
Unknown | NPOC | Vial_41 | Untitled | 2025-06-11 05:37 | 1 | 1 | NPOC | 37.26 | 34.47 | 0 | 0.6576 | 1.91 | 41 | 1 | 100 | 1
Unknown | NPOC | Vial_41 | Untitled | 2025-06-11 05:40 | 1 | 2 | NPOC | 34.93 | 34.47 | 0 | 0.6576 | 1.91 | 41 | 0 | 100 | 1
Unknown | NPOC | Vial_41 | Untitled | 2025-06-11 05:43 | 1 | 3 | NPOC | 34 | 34.47 | 0 | 0.6576 | 1.91 | 41 | 0 | 100 | 1
Unknown | NPOC | Vial_42 | Untitled | 2025-06-11 05:54 | 1 | 1 | NPOC | 33.7 | 30.75 | 0 | 0.4313 | 1.4 | 42 | 1 | 100 | 1
Unknown | NPOC | Vial_42 | Untitled | 2025-06-11 05:57 | 1 | 2 | NPOC | 31.05 | 30.75 | 0 | 0.4313 | 1.4 | 42 | 0 | 100 | 1
Unknown | NPOC | Vial_42 | Untitled | 2025-06-11 06:00 | 1 | 3 | NPOC | 30.44 | 30.75 | 0 | 0.4313 | 1.4 | 42 | 0 | 100 | 1
Unknown | NPOC | Vial_12_QC4 | Untitled | 2025-06-11 06:12 | 1 | 1 | NPOC | 75.44 | 71.41 | 0 | 0.4243 | 0.59 | 12 | 1 | 100 | 1
Unknown | NPOC | Vial_12_QC4 | Untitled | 2025-06-11 06:15 | 1 | 2 | NPOC | 71.71 | 71.41 | 0 | 0.4243 | 0.59 | 12 | 0 | 100 | 1
Unknown | NPOC | Vial_12_QC4 | Untitled | 2025-06-11 06:18 | 1 | 3 | NPOC | 71.11 | 71.41 | 0 | 0.4243 | 0.59 | 12 | 0 | 100 | 1
Unknown | NPOC | Vial_17_blank4 | Untitled | 2025-06-11 06:28 | 1 | 1 | NPOC | 1.167 | 0.5535 | 0 | 0.1105 | 19.97 | 17 | 1 | 100 | 1
Unknown | NPOC | Vial_17_blank4 | Untitled | 2025-06-11 06:30 | 1 | 2 | NPOC | 0.6317 | 0.5535 | 0 | 0.1105 | 19.97 | 17 | 0 | 100 | 1
Unknown | NPOC | Vial_17_blank4 | Untitled | 2025-06-11 06:32 | 1 | 3 | NPOC | 0.4754 | 0.5535 | 0 | 0.1105 | 19.97 | 17 | 0 | 100 | 1
Unknown | NPOC | Vial_49_VDOW_0_REP | Untitled | 2025-06-11 06:43 | 1 | 1 | NPOC | 29.71 | 29.62 | 0 | 0.1273 | 0.43 | 49 | 0 | 100 | 1
Unknown | NPOC | Vial_49_VDOW_0_REP | Untitled | 2025-06-11 06:46 | 1 | 2 | NPOC | 29.53 | 29.62 | 0 | 0.1273 | 0.43 | 49 | 0 | 100 | 1
Unknown | NPOC | Vial_50_VDOW_10_REP | Untitled | 2025-06-11 06:57 | 1 | 1 | NPOC | 98.37 | 97.31 | 0 | 1.492 | 1.53 | 50 | 0 | 100 | 1
Unknown | NPOC | Vial_50_VDOW_10_REP | Untitled | 2025-06-11 07:01 | 1 | 2 | NPOC | 96.26 | 97.31 | 0 | 1.492 | 1.53 | 50 | 0 | 100 | 1
Unknown | NPOC | nanopure | Untitled | 2025-06-11 07:11 | 1 | 1 | NPOC | 5.232 | 1.485 | 0 | 0.7126 | 47.98 | 1 | 1 | 100 | 1
Unknown | NPOC | nanopure | Untitled | 2025-06-11 07:13 | 1 | 2 | NPOC | 1.989 | 1.485 | 0 | 0.7126 | 47.98 | 1 | 0 | 100 | 1
Unknown | NPOC | nanopure | Untitled | 2025-06-11 07:15 | 1 | 3 | NPOC | 0.9813 | 1.485 | 0 | 0.7126 | 47.98 | 1 | 0 | 100 | 1
Unknown | NPOC | Vial_51 | Untitled | 2025-06-11 07:28 | 1 | 1 | NPOC | 1272 | 1268 | 0 | 6.364 | 0.5 | 51 | 0 | 100 | 1
Unknown | NPOC | Vial_51 | Untitled | 2025-06-11 07:32 | 1 | 2 | NPOC | 1263 | 1268 | 0 | 6.364 | 0.5 | 51 | 0 | 100 | 1
Unknown | NPOC | Vial_52 | Untitled | 2025-06-11 07:44 | 1 | 1 | NPOC | 202.3 | 201.4 | 0 | 1.344 | 0.67 | 52 | 0 | 100 | 1
Unknown | NPOC | Vial_52 | Untitled | 2025-06-11 07:47 | 1 | 2 | NPOC | 200.4 | 201.4 | 0 | 1.344 | 0.67 | 52 | 0 | 100 | 1
Unknown | NPOC | Vial_53 | Untitled | 2025-06-11 07:59 | 1 | 1 | NPOC | 142.2 | 140.4 | 0 | 2.616 | 1.86 | 53 | 0 | 100 | 1
Unknown | NPOC | Vial_53 | Untitled | 2025-06-11 08:03 | 1 | 2 | NPOC | 138.5 | 140.4 | 0 | 2.616 | 1.86 | 53 | 0 | 100 | 1
Unknown | NPOC | Vial_54 | Untitled | 2025-06-11 08:14 | 1 | 1 | NPOC | 58.59 | 58.05 | 0 | 0.7637 | 1.32 | 54 | 0 | 100 | 1
Unknown | NPOC | Vial_54 | Untitled | 2025-06-11 08:18 | 1 | 2 | NPOC | 57.51 | 58.05 | 0 | 0.7637 | 1.32 | 54 | 0 | 100 | 1
Unknown | NPOC | Vial_55 | Untitled | 2025-06-11 08:29 | 1 | 1 | NPOC | 33.43 | 33.04 | 0 | 0.5515 | 1.67 | 55 | 0 | 100 | 1
Unknown | NPOC | Vial_55 | Untitled | 2025-06-11 08:32 | 1 | 2 | NPOC | 32.65 | 33.04 | 0 | 0.5515 | 1.67 | 55 | 0 | 100 | 1
Unknown | NPOC | Vial_56 | Untitled | 2025-06-11 08:43 | 1 | 1 | NPOC | 19.12 | 18.48 | 0 | 0.2828 | 1.53 | 56 | 1 | 100 | 1
Unknown | NPOC | Vial_56 | Untitled | 2025-06-11 08:46 | 1 | 2 | NPOC | 18.28 | 18.48 | 0 | 0.2828 | 1.53 | 56 | 0 | 100 | 1
Unknown | NPOC | Vial_56 | Untitled | 2025-06-11 08:49 | 1 | 3 | NPOC | 18.68 | 18.48 | 0 | 0.2828 | 1.53 | 56 | 0 | 100 | 1
Unknown | NPOC | Vial_57 | Untitled | 2025-06-11 08:59 | 1 | 1 | NPOC | 13.55 | 12.57 | 0 | 0.198 | 1.58 | 57 | 1 | 100 | 1
Unknown | NPOC | Vial_57 | Untitled | 2025-06-11 09:02 | 1 | 2 | NPOC | 12.43 | 12.57 | 0 | 0.198 | 1.58 | 57 | 0 | 100 | 1
Unknown | NPOC | Vial_57 | Untitled | 2025-06-11 09:04 | 1 | 3 | NPOC | 12.71 | 12.57 | 0 | 0.198 | 1.58 | 57 | 0 | 100 | 1
Unknown | NPOC | Vial_58 | Untitled | 2025-06-11 09:14 | 1 | 1 | NPOC | 6.187 | 5.795 | 0 | 0.04031 | 0.7 | 58 | 1 | 100 | 1
Unknown | NPOC | Vial_58 | Untitled | 2025-06-11 09:17 | 1 | 2 | NPOC | 5.766 | 5.795 | 0 | 0.04031 | 0.7 | 58 | 0 | 100 | 1
Unknown | NPOC | Vial_58 | Untitled | 2025-06-11 09:19 | 1 | 3 | NPOC | 5.823 | 5.795 | 0 | 0.04031 | 0.7 | 58 | 0 | 100 | 1
Unknown | NPOC | Vial_59 | Untitled | 2025-06-11 09:29 | 1 | 1 | NPOC | 3.366 | 3.197 | 0 | 0.00636 | 0.2 | 59 | 1 | 100 | 1
Unknown | NPOC | Vial_59 | Untitled | 2025-06-11 09:31 | 1 | 2 | NPOC | 3.201 | 3.197 | 0 | 0.00636 | 0.2 | 59 | 0 | 100 | 1
Unknown | NPOC | Vial_59 | Untitled | 2025-06-11 09:34 | 1 | 3 | NPOC | 3.192 | 3.197 | 0 | 0.00636 | 0.2 | 59 | 0 | 100 | 1
Unknown | NPOC | Vial_60 | Untitled | 2025-06-11 09:44 | 1 | 1 | NPOC | 9.288 | 8.958 | 0 | 0.07212 | 0.81 | 60 | 1 | 100 | 1
Unknown | NPOC | Vial_60 | Untitled | 2025-06-11 09:47 | 1 | 2 | NPOC | 9.009 | 8.958 | 0 | 0.07212 | 0.81 | 60 | 0 | 100 | 1
Unknown | NPOC | Vial_60 | Untitled | 2025-06-11 09:49 | 1 | 3 | NPOC | 8.907 | 8.958 | 0 | 0.07212 | 0.81 | 60 | 0 | 100 | 1
Unknown | NPOC | Vial_61 | Untitled | 2025-06-11 09:59 | 1 | 1 | NPOC | 5.311 | 5.064 | 0 | 0.07566 | 1.49 | 61 | 1 | 100 | 1
Unknown | NPOC | Vial_61 | Untitled | 2025-06-11 10:01 | 1 | 2 | NPOC | 5.01 | 5.064 | 0 | 0.07566 | 1.49 | 61 | 0 | 100 | 1
Unknown | NPOC | Vial_61 | Untitled | 2025-06-11 10:04 | 1 | 3 | NPOC | 5.117 | 5.064 | 0 | 0.07566 | 1.49 | 61 | 0 | 100 | 1
Unknown | NPOC | Vial_62 | Untitled | 2025-06-11 10:14 | 1 | 1 | NPOC | 2.31 | 2.276 | 0 | 0.04808 | 2.11 | 62 | 0 | 100 | 1
Unknown | NPOC | Vial_62 | Untitled | 2025-06-11 10:16 | 1 | 2 | NPOC | 2.242 | 2.276 | 0 | 0.04808 | 2.11 | 62 | 0 | 100 | 1
Unknown | NPOC | Vial_63 | Untitled | 2025-06-11 10:26 | 1 | 1 | NPOC | 1.46 | 1.397 | 0 | 0.0898 | 6.43 | 63 | 0 | 100 | 1
Unknown | NPOC | Vial_63 | Untitled | 2025-06-11 10:28 | 1 | 2 | NPOC | 1.333 | 1.397 | 0 | 0.0898 | 6.43 | 63 | 0 | 100 | 1
Unknown | NPOC | Vial_64 | Untitled | 2025-06-11 10:38 | 1 | 1 | NPOC | 1.407 | 0.9338 | 0 | 0.04087 | 4.38 | 64 | 1 | 100 | 1
Unknown | NPOC | Vial_64 | Untitled | 2025-06-11 10:40 | 1 | 2 | NPOC | 0.9627 | 0.9338 | 0 | 0.04087 | 4.38 | 64 | 0 | 100 | 1
Unknown | NPOC | Vial_64 | Untitled | 2025-06-11 10:42 | 1 | 3 | NPOC | 0.9049 | 0.9338 | 0 | 0.04087 | 4.38 | 64 | 0 | 100 | 1
Unknown | NPOC | Vial_65 | Untitled | 2025-06-11 10:52 | 1 | 1 | NPOC | 0.8998 | 0.7123 | 0 | 0.02454 | 3.44 | 65 | 1 | 100 | 1
Unknown | NPOC | Vial_65 | Untitled | 2025-06-11 10:54 | 1 | 2 | NPOC | 0.7296 | 0.7123 | 0 | 0.02454 | 3.44 | 65 | 0 | 100 | 1
Unknown | NPOC | Vial_65 | Untitled | 2025-06-11 10:56 | 1 | 3 | NPOC | 0.6949 | 0.7123 | 0 | 0.02454 | 3.44 | 65 | 0 | 100 | 1
Unknown | NPOC | Vial_66 | Untitled | 2025-06-11 11:06 | 1 | 1 | NPOC | 0.9007 | 0.5763 | 0 | 0.03352 | 5.82 | 66 | 1 | 100 | 1
Unknown | NPOC | Vial_66 | Untitled | 2025-06-11 11:08 | 1 | 2 | NPOC | 0.5526 | 0.5763 | 0 | 0.03352 | 5.82 | 66 | 0 | 100 | 1
Unknown | NPOC | Vial_66 | Untitled | 2025-06-11 11:10 | 1 | 3 | NPOC | 0.6 | 0.5763 | 0 | 0.03352 | 5.82 | 66 | 0 | 100 | 1
Unknown | NPOC | Vial_67 | Untitled | 2025-06-11 11:20 | 1 | 1 | NPOC | 1.094 | 0.5625 | 0 | 0.01768 | 3.14 | 67 | 1 | 100 | 1
Unknown | NPOC | Vial_67 | Untitled | 2025-06-11 11:22 | 1 | 2 | NPOC | 0.55 | 0.5625 | 0 | 0.01768 | 3.14 | 67 | 0 | 100 | 1
Unknown | NPOC | Vial_67 | Untitled | 2025-06-11 11:24 | 1 | 3 | NPOC | 0.575 | 0.5625 | 0 | 0.01768 | 3.14 | 67 | 0 | 100 | 1
Unknown | NPOC | Vial_68 | Untitled | 2025-06-11 11:34 | 1 | 1 | NPOC | 0.8897 | 0.8285 | 0 | 0.08662 | 10.46 | 68 | 0 | 100 | 1
Unknown | NPOC | Vial_68 | Untitled | 2025-06-11 11:36 | 1 | 2 | NPOC | 0.7672 | 0.8285 | 0 | 0.08662 | 10.46 | 68 | 0 | 100 | 1
Unknown | NPOC | Vial_69 | Untitled | 2025-06-11 11:46 | 1 | 1 | NPOC | 1.165 | 1.136 | 0 | 0.04172 | 3.67 | 69 | 0 | 100 | 1
Unknown | NPOC | Vial_69 | Untitled | 2025-06-11 11:48 | 1 | 2 | NPOC | 1.106 | 1.136 | 0 | 0.04172 | 3.67 | 69 | 0 | 100 | 1
Unknown | NPOC | Vial_70 | Untitled | 2025-06-11 11:58 | 1 | 1 | NPOC | 3.244 | 3.223 | 0 | 0.0297 | 0.92 | 70 | 0 | 100 | 1
Unknown | NPOC | Vial_70 | Untitled | 2025-06-11 12:00 | 1 | 2 | NPOC | 3.202 | 3.223 | 0 | 0.0297 | 0.92 | 70 | 0 | 100 | 1
Unknown | NPOC | Vial_13_QC5 | Untitled | 2025-06-11 12:12 | 1 | 1 | NPOC | 72.89 | 72.24 | 0 | 0.9192 | 1.27 | 13 | 0 | 100 | 1
Unknown | NPOC | Vial_13_QC5 | Untitled | 2025-06-11 12:15 | 1 | 2 | NPOC | 71.59 | 72.24 | 0 | 0.9192 | 1.27 | 13 | 0 | 100 | 1
Unknown | NPOC | Vial_18_blank5 | Untitled | 2025-06-11 12:25 | 1 | 1 | NPOC | 0.4519 | 0.3891 | 0 | 0.08881 | 22.83 | 18 | 0 | 100 | 1
Unknown | NPOC | Vial_18_blank5 | Untitled | 2025-06-11 12:27 | 1 | 2 | NPOC | 0.3263 | 0.3891 | 0 | 0.08881 | 22.83 | 18 | 0 | 100 | 1
Unknown | NPOC | 0position | Untitled | 2025-06-11 12:35 | 1 | 1 | NPOC | 0.831 | 0.8876 | 0 | 0.07997 | 9.01 | 0 | 0 | 100 | 1
Unknown | NPOC | 0position | Untitled | 2025-06-11 12:38 | 1 | 2 | NPOC | 0.5534 | 0.8876 | 0 | 0.07997 | 9.01 | 0 | 1 | 100 | 1
Unknown | NPOC | 0position | Untitled | 2025-06-11 12:40 | 1 | 3 | NPOC | 0.9441 | 0.8876 | 0 | 0.07997 | 9.01 | 0 | 0 | 100 | 1

## working_data
Sample Name | Analysis(Inj.) | Mean Area | Conc. | Vial
0position | NPOC | 0.8424 | 0 | 0
nanopure | NPOC | 0.4632 | 0 | 1
NPOC_CAL | NPOC | 137.3 | 20 | 2
NPOC_CAL | NPOC | 69.06 | 10 | 3
NPOC_CAL | NPOC | 34.2 | 5 | 4
NPOC_CAL | NPOC | 14.19 | 2 | 5
NPOC_CAL | NPOC | 7.179 | 1 | 6
NPOC_CAL | NPOC | 3.66 | 0.5 | 7
NPOC_CAL | NPOC | 0.3301 | 0 | 8
Vial_19 | NPOC | 88.43 | 0 | 19
Vial_20 | NPOC | 56.79 | 0 | 20
Vial_21 | NPOC | 38.2 | 0 | 21
Vial_22 | NPOC | 30.55 | 0 | 22
Vial_23 | NPOC | 28.28 | 0 | 23
Vial_24 | NPOC | 23.82 | 0 | 24
Vial_9_QC1 | NPOC | 69.55 | 0 | 9
Vial_14_blank1 | NPOC | 0.5861 | 0 | 14
Vial_43_SLOC_0_REP | NPOC | 22.76 | 0 | 43
Vial_44_SLOC_10_REP | NPOC | 91.35 | 0 | 44
Vial_25 | NPOC | 82 | 0 | 25
Vial_26 | NPOC | 47.88 | 0 | 26
Vial_27 | NPOC | 28.41 | 0 | 27
Vial_28 | NPOC | 20.77 | 0 | 28
Vial_29 | NPOC | 16.89 | 0 | 29
Vial_30 | NPOC | 13.77 | 0 | 30
Vial_10_QC2 | NPOC | 71.44 | 0 | 10
Vial_15_blank2 | NPOC | 0.4273 | 0 | 15
Vial_45_SLOW_0_REP | NPOC | 12.7 | 0 | 45
Vial_46_SLOW_10_REP | NPOC | 80.74 | 0 | 46
Vial_31 | NPOC | 101.5 | 0 | 31
Vial_32 | NPOC | 66.84 | 0 | 32
Vial_33 | NPOC | 46.77 | 0 | 33
Vial_34 | NPOC | 40.36 | 0 | 34
Vial_35 | NPOC | 36.56 | 0 | 35
Vial_36 | NPOC | 33.78 | 0 | 36
Vial_11_QC3 | NPOC | 71.33 | 0 | 11
Vial_16_blank3 | NPOC | 0.6499 | 0 | 16
Vial_47_VDOS_0_REP | NPOC | 31.76 | 0 | 47
Vial_48_VDOS_10_REP | NPOC | 100.4 | 0 | 48
Vial_37 | NPOC | 95.62 | 0 | 37
Vial_38 | NPOC | 63.85 | 0 | 38
Vial_39 | NPOC | 44.45 | 0 | 39
Vial_40 | NPOC | 38.48 | 0 | 40
Vial_41 | NPOC | 34.47 | 0 | 41
Vial_42 | NPOC | 30.75 | 0 | 42
Vial_12_QC4 | NPOC | 71.41 | 0 | 12
Vial_17_blank4 | NPOC | 0.5535 | 0 | 17
Vial_49_VDOW_0_REP | NPOC | 29.62 | 0 | 49
Vial_50_VDOW_10_REP | NPOC | 97.31 | 0 | 50
nanopure | NPOC | 1.485 | 0 | 1
Vial_51 | NPOC | 1268 | 0 | 51
Vial_52 | NPOC | 201.4 | 0 | 52
Vial_53 | NPOC | 140.4 | 0 | 53
Vial_54 | NPOC | 58.05 | 0 | 54
Vial_55 | NPOC | 33.04 | 0 | 55
Vial_56 | NPOC | 18.48 | 0 | 56
Vial_57 | NPOC | 12.57 | 0 | 57
Vial_58 | NPOC | 5.795 | 0 | 58
Vial_59 | NPOC | 3.197 | 0 | 59
Vial_60 | NPOC | 8.958 | 0 | 60
Vial_61 | NPOC | 5.064 | 0 | 61
Vial_62 | NPOC | 2.276 | 0 | 62
Vial_63 | NPOC | 1.397 | 0 | 63
Vial_64 | NPOC | 0.9338 | 0 | 64
Vial_65 | NPOC | 0.7123 | 0 | 65
Vial_66 | NPOC | 0.5763 | 0 | 66
Vial_67 | NPOC | 0.5625 | 0 | 67
Vial_68 | NPOC | 0.8285 | 0 | 68
Vial_69 | NPOC | 1.136 | 0 | 69
Vial_70 | NPOC | 3.223 | 0 | 70
Vial_13_QC5 | NPOC | 72.24 | 0 | 13
Vial_18_blank5 | NPOC | 0.3891 | 0 | 18
0position | NPOC | 0.8876 | 0 | 0

## cal_curves
Complete Calibration Curves |  |  | Calibration summary from instrument
NPOC beginning |  |  | Analysis | r2 | Standard Error | Intercept | Slope
Sample ID | Mean Area | Prep Conc (g/mL) | NPOC
Blank | 0.3301 | 0 | TN
Std 1 | 137.3 | 20
Std 2 | 69.06 | 10 | NPOC beginning | [8 empty] | TN beginning
Std 3 | 34.2 | 5 | SUMMARY OUTPUT | [8 empty] | SUMMARY OUTPUT
Std 4 | 14.19 | 2
Std 5 | 7.179 | 1 | Regression Statistics
Std 6 | 3.66 | 0.5 | Multiple R | 0.99999
|  |  | R Square | 0.99998
|  |  | Adjusted R Square | 0.99998
|  |  | Standard Error | 0.21416
|  |  | Observations | 7
ANOVA
|  |  |  | df | SS | MS | F | Significance F
|  |  | Regression | 1 | 14954.5304 | 14954.5304 | 326049.726 | 3.1267e-13
|  |  | Residual | 5 | 0.22933 | 0.045866
|  |  | Total | 6 | 14954.7598
|  |  |  | Coefficients | Standard Error | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
|  |  | Intercept | 0.3012 | 0.10444 | 2.8839 | 0.034432 | 0.03272 | 0.56968 | 0.03272 | 0.56968
|  |  | Prep Conc (g/mL) | 6.8522 | 0.012 | 571.0076 | 3.1267e-13 | 6.8214 | 6.8831 | 6.8214 | 6.8831
RESIDUAL OUTPUT
|  |  | Observation | Predicted Mean Area | Residuals | Standard Residuals
|  |  | 1 | 0.3012 | 0.0289 | 0.14782
|  |  | 2 | 137.3457 | -0.045718 | -0.23385
|  |  | 3 | 68.8235 | 0.23654 | 1.2099
|  |  | 4 | 34.5623 | -0.36233 | -1.8533
|  |  | 5 | 14.0057 | 0.18435 | 0.94294
|  |  | 6 | 7.1534 | 0.025574 | 0.13081
|  |  | 7 | 3.7273 | -0.067313 | -0.34431
[13 empty] | Predicted Mean Area | Residuals | Standard Residuals
[13 empty] | -2.9341 | 2.9341 | 0.40251
[13 empty] | 263.6162 | 5.7838 | 0.79345
[13 empty] | 130.341 | -12.741 | -1.7479
[13 empty] | 50.376 | 1.844 | 0.25297
[13 empty] | 23.7209 | 2.1791 | 0.29894

## calc_data
NPOC (non-purgeable organic carbon) | [6 empty] | Final values
Sample Vial | Sample | Mean Area | Conc (mg/L) | Matrix Spike | (Spiked - Unspiked) (hoping for Column E) | % Recovery (hoping for 1) |  | Conc (mg/L) | Matrix Spike | (Spiked - Unspiked) | % Correct from expected
Vial_19 | SLOC 10 mg/L matrix spike | 88.43 | 12.8613 | 10 | 9.4291 | 0.94291 | SLOC 10 mg/L matrix spike | 12.8613 | 10 | 9.4291 | 94.29%
Vial_20 | SLOC 5 mg/L matrix spike | 56.79 | 8.2439 | 5 | 4.8116 | 0.96232 | SLOC 5 mg/L matrix spike | 8.2439 | 5 | 4.8116 | 96.23%
Vial_21 | SLOC 2 mg/L matrix spike | 38.2 | 5.5309 | 2 | 2.0986 | 1.0493 | SLOC 2 mg/L matrix spike | 5.5309 | 2 | 2.0986 | 104.9%
Vial_22 | SLOC 1 mg/L matrix spike | 30.55 | 4.4144 | 1 | 0.98216 | 0.98216 | SLOC 1 mg/L matrix spike | 4.4144 | 1 | 0.98216 | 98.22%
Vial_23 | SLOC 0.5 mg/L matrix spike | 28.28 | 4.0832 | 0.5 | 0.65088 | 1.3018 | SLOC 0.5 mg/L matrix spike | 4.0832 | 0.5 | 0.65088 | 130.2%
Vial_24 | SLOC plain matrix | 23.82 | 3.4323 | 0 |  |  | SLOC plain matrix | 3.4323 | 0
Vial_9_QC1 |  | 69.55 | 10.106 |  |  |  | SLOW 10 mg/L matrix spike | 11.923 | 10 | 9.9573 | 99.57%
Vial_14_blank1 |  | 0.5861 | 0.041578 |  |  |  | SLOW 5 mg/L matrix spike | 6.9436 | 5 | 4.9779 | 99.56%
Vial_43_SLOC_0_REP |  | 22.76 | 3.2776 | 0 |  |  | SLOW 2 mg/L matrix spike | 4.1021 | 2 | 2.1365 | 106.8%
Vial_44_SLOC_10_REP |  | 91.35 | 13.2875 | 10 |  |  | SLOW 1 mg/L matrix spike | 2.9872 | 1 | 1.0216 | 102.2%
Vial_25 | SLOW 10 mg/L matrix spike | 82 | 11.923 | 10 | 9.9573 | 0.99573 | SLOW 0.5 mg/L matrix spike | 2.4209 | 0.5 | 0.45533 | 91.07%
Vial_26 | SLOW 5 mg/L matrix spike | 47.88 | 6.9436 | 5 | 4.9779 | 0.99559 | SLOW plain matrix | 1.9656 | 0
Vial_27 | SLOW 2 mg/L matrix spike | 28.41 | 4.1021 | 2 | 2.1365 | 1.0683 | VDOS 10 mg/L matrix spike | 14.7687 | 10 | 9.8829 | 98.83%
Vial_28 | SLOW 1 mg/L matrix spike | 20.77 | 2.9872 | 1 | 1.0216 | 1.0216 | VDOS 5 mg/L matrix spike | 9.7105 | 5 | 4.8247 | 96.49%
Vial_29 | SLOW 0.5 mg/L matrix spike | 16.89 | 2.4209 | 0.5 | 0.45533 | 0.91065 | VDOS 2 mg/L matrix spike | 6.7816 | 2 | 1.8957 | 94.79%
Vial_30 | SLOW plain matrix | 13.77 | 1.9656 | 0 |  |  | VDOS 1 mg/L matrix spike | 5.8461 | 1 | 0.96027 | 96.03%
Vial_10_QC2 |  | 71.44 | 10.3819 |  |  |  | VDOS 0.5 mg/L matrix spike | 5.2915 | 0.5 | 0.40571 | 81.14%
Vial_15_blank2 |  | 0.4273 | 0.018403 |  |  |  | VDOS plain matrix | 4.8858 | 0
Vial_45_SLOW_0_REP |  | 12.7 | 1.8095 | 0 |  |  | VDOW 10 mg/L matrix spike | 13.9106 | 10 | 9.467 | 94.67%
Vial_46_SLOW_10_REP |  | 80.74 | 11.7391 | 10 |  |  | VDOW 5 mg/L matrix spike | 9.2742 | 5 | 4.8305 | 96.61%
Vial_31 | VDOS 10 mg/L matrix spike | 101.5 | 14.7687 | 10 | 9.8829 | 0.98829 | VDOW 2 mg/L matrix spike | 6.443 | 2 | 1.9994 | 99.97%
Vial_32 | VDOS 5 mg/L matrix spike | 66.84 | 9.7105 | 5 | 4.8247 | 0.96494 | VDOW 1 mg/L matrix spike | 5.5717 | 1 | 1.1281 | 112.8%
Vial_33 | VDOS 2 mg/L matrix spike | 46.77 | 6.7816 | 2 | 1.8957 | 0.94787 | VDOW 0.5 mg/L matrix spike | 4.9865 | 0.5 | 0.54289 | 108.6%
Vial_34 | VDOS 1 mg/L matrix spike | 40.36 | 5.8461 | 1 | 0.96027 | 0.96027 | VDOW plain matrix | 4.4436 | 0
Vial_35 | VDOS 0.5 mg/L matrix spike | 36.56 | 5.2915 | 0.5 | 0.40571 | 0.81142
Vial_36 | VDOS plain matrix | 33.78 | 4.8858 | 0
Vial_11_QC3 |  | 71.33 | 10.3658
Vial_16_blank3 |  | 0.6499 | 0.050889
Vial_47_VDOS_0_REP |  | 31.76 | 4.591 | 0
Vial_48_VDOS_10_REP |  | 100.4 | 14.6082 | 10
Vial_37 | VDOW 10 mg/L matrix spike | 95.62 | 13.9106 | 10 | 9.467 | 0.9467
Vial_38 | VDOW 5 mg/L matrix spike | 63.85 | 9.2742 | 5 | 4.8305 | 0.96611
Vial_39 | VDOW 2 mg/L matrix spike | 44.45 | 6.443 | 2 | 1.9994 | 0.99968
Vial_40 | VDOW 1 mg/L matrix spike | 38.48 | 5.5717 | 1 | 1.1281 | 1.1281
Vial_41 | VDOW 0.5 mg/L matrix spike | 34.47 | 4.9865 | 0.5 | 0.54289 | 1.0858
Vial_42 | VDOW plain matrix | 30.75 | 4.4436 | 0
Vial_12_QC4 |  | 71.41 | 10.3775
Vial_17_blank4 |  | 0.5535 | 0.03682
Vial_49_VDOW_0_REP |  | 29.62 | 4.2787 | 0
Vial_50_VDOW_10_REP |  | 97.31 | 14.1573 | 10
nanopure |  | 1.485 | 0.17276 |  |  |  | Sample | Carbon (mg/L) |  |  |  | Sample | Carbon (mg/L)
Vial_51 | filter test: sample 1 | 1268 | 185.0054 |  |  |  | 5 mL to wet + 5 mL pushed through | 185.0054 |  |  |  | 10-15 mL | 29.348
Vial_52 | filter test: sample 2 | 201.4 | 29.348 |  |  |  | 10-15 mL | 29.348 |  |  |  | 15-20 mL | 20.4457
Vial_53 | filter test: sample 3 | 140.4 | 20.4457 |  |  |  | 15-20 mL | 20.4457 |  |  |  | 20-25 mL | 8.4277
Vial_54 | filter test: sample 4 | 58.05 | 8.4277 |  |  |  | 20-25 mL | 8.4277 |  |  |  | 25-30 mL | 4.7778
Vial_55 | filter test: sample 5 | 33.04 | 4.7778 |  |  |  | 25-30 mL | 4.7778 |  |  |  | 30-35 mL | 2.653
Vial_56 | filter test: sample 6 | 18.48 | 2.653 |  |  |  | 30-35 mL | 2.653 |  |  |  | 35-40 mL | 1.7905
Vial_57 | filter test: sample 7 | 12.57 | 1.7905 |  |  |  | 35-40 mL | 1.7905 |  |  |  | 40-45 mL | 0.80175
Vial_58 | filter test: sample 8 | 5.795 | 0.80175 |  |  |  | 40-45 mL | 0.80175 |  |  |  | 45-50 mL | 0.42261
Vial_59 | filter test: sample 9 | 3.197 | 0.42261 |  |  |  | 45-50 mL | 0.42261 |  |  |  | refill syringe, 50-55 mL | 1.2634
Vial_60 | filter test: sample 10 | 8.958 | 1.2634 |  |  |  | refill syringe, 50-55 mL | 1.2634 |  |  |  | 55-60 mL | 0.69507
Vial_61 | filter test: sample 11 | 5.064 | 0.69507 |  |  |  | 55-60 mL | 0.69507 |  |  |  | 60-65 mL | 0.2882
Vial_62 | filter test: sample 12 | 2.276 | 0.2882 |  |  |  | 60-65 mL | 0.2882 |  |  |  | 65-70 mL | 0.15992
Vial_63 | filter test: sample 13 | 1.397 | 0.15992 |  |  |  | 65-70 mL | 0.15992 |  |  |  | 70-75 mL | 0.09232
Vial_64 | filter test: sample 14 | 0.9338 | 0.09232 |  |  |  | 70-75 mL | 0.09232 |  |  |  | 75-80 mL | 0.059995
Vial_65 | filter test: sample 15 | 0.7123 | 0.059995 |  |  |  | 75-80 mL | 0.059995 |  |  |  | 80-85 mL | 0.040147
Vial_66 | filter test: sample 16 | 0.5763 | 0.040147 |  |  |  | 80-85 mL | 0.040147 |  |  |  | 85-90 mL | 0.038134
Vial_67 | filter test: sample 17 | 0.5625 | 0.038134 |  |  |  | 85-90 mL | 0.038134 |  |  |  | 90-95 mL | 0.076953
Vial_68 | filter test: sample 18 | 0.8285 | 0.076953 |  |  |  | 90-95 mL | 0.076953 |  |  |  | 95-100 mL | 0.12183
Vial_69 | filter test: sample 19 | 1.136 | 0.12183 |  |  |  | 95-100 mL | 0.12183 |  |  |  | refill syringe, 100-105 mL | 0.4264
Vial_70 | filter test: sample 20 | 3.223 | 0.4264 |  |  |  | refill syringe, 100-105 mL | 0.4264
Vial_13_QC5 |  | 72.24 | 10.4986
Vial_18_blank5 |  | 0.3891 | 0.012828

## qc_rep_blanks
NPOC QC
Prepared Concentration = |  |  | 10 | mg/L
QC # | Mean Area | Measured Conc. | Error | Error Avg
QC 1 | 69.55 | 10.106 | 1.06% | 3.46%
QC 2 | 71.44 | 10.3819 | 3.819%
QC 3 | 71.33 | 10.3658 | 3.658%
QC 4 | 71.41 | 10.3775 | 3.775%
QC 5 | 72.24 | 10.4986 | 4.986%
Replicates (Different Vial)
Sample Name | Sample ID | NPOC | NPOC Avg | NPOC % Errors
SLOC plain matrix | Vial_24 | 3.4323 | 3.3549 | 4.507%
| Vial_43_SLOC_0_REP | 3.2776
SLOC 10 mg/L matrix spike | Vial_19 | 12.8613 | 13.0744 | -3.313%
| Vial_44_SLOC_10_REP | 13.2875
SLOW plain matrix | Vial_30 | 1.9656 | 1.8875 | 7.944%
| Vial_45_SLOW_0_REP | 1.8095
SLOW 10 mg/L matrix spike | Vial_25 | 11.923 | 11.831 | 1.542%
| Vial_46_SLOW_10_REP | 11.7391
VDOS plain matrix | Vial_36 | 4.8858 | 4.7384 | 6.034%
| Vial_47_VDOS_0_REP | 4.591
VDOS 10 mg/L matrix spike | Vial_31 | 14.7687 | 14.6885 | 1.087%
| Vial_48_VDOS_10_REP | 14.6082
VDOW plain matrix | Vial_42 | 4.4436 | 4.3612 | 3.711%
| Vial_49_VDOW_0_REP | 4.2787
VDOW 10 mg/L matrix spike | Vial_37 | 13.9106 | 14.034 | -1.773%
| Vial_50_VDOW_10_REP | 14.1573
Blanks
Blank # | Mean Area | Conc (mg/L)
| NPOC | NPOC
Vial_14_blank1 | 0.5861 | 0.041578
Vial_15_blank2 | 0.4273 | 0.018403
Vial_16_blank3 | 0.6499 | 0.050889
Vial_17_blank4 | 0.5535 | 0.03682
Vial_18_blank5 | 0.3891 | 0.012828

## calculations
Vial | [5 empty] | Cal Curve Standard (NPOC/TN) |  | NPOC Stock Mass | QC |  |  | NPOC mass
# | Sample Name | Info | Sample mL | Dilution mL | Total mL | Stock Conc (ppm) | 1000 |  | Stock Conc (ppm) | 1000 | QC 1 | 0.2462
1 | nanopure flush |  |  |  |  | Vol Flask Final (mL) | 100 |  | Vol Flask Final (mL) | 25 | QC 2 | 0.2513
2 | 20 mg/L NPOC |  |  |  |  | NPOC Desired Conc (mg/L) | 20 |  | NPOC Desired Conc (mg/L) | 10 | QC 3 | 0.2517
3 | 10 mg/L NPOC | [9 empty] | QC 3 | 0.2523
4 | 5 mg/L NPOC |  |  |  |  | NPOC Vol Needed (mL) | 2 | 1.9464 | NPOC Vol Needed (mL) | 0.25 | QC 5 | 0.2514
5 | 2 mg/L NPOC
6 | 1 mg/L NPOC
7 | 0.5 mg/L NPOC |  |  |  |  | Stock Conc (ppm) | 1000
8 | nanopure |  |  |  |  | Vol Flask Final (mL) | 100
9 | QC 1 (10 mg-C/L) |  |  |  |  | NPOC Desired Conc (mg/L) | 10
10 | QC 2 (10 mg-C/L)
11 | QC 3 (10 mg-C/L) |  |  |  |  | NPOC Vol Needed (mL) | 1 | 0.9767
12 | QC 4 (10 mg-C/L)
13 | QC 5 (10 mg-C/L)
14 | blank 1 (nanopure) |  |  |  |  | Stock Conc (ppm) | 1000
15 | blank 2 (nanopure) |  |  |  |  | Vol Flask Final (mL) | 100
16 | blank 3 (nanopure) |  |  |  |  | NPOC Desired Conc (mg/L) | 5
17 | blank 4 (nanopure)
18 | blank 5 (nanopure) |  |  |  |  | NPOC Vol Needed (mL) | 0.5 | 0.4898
19 | SLOC 10 mg/L matrix spike | filter 2
20 | SLOC 5 mg/L matrix spike | filter 2
21 | SLOC 2 mg/L matrix spike | filter 2 | rinsed with 50 mL plain SLOC |  |  | Stock Conc (ppm) | 1000
22 | SLOC 1 mg/L matrix spike | filter 1 |  |  |  | Vol Flask Final (mL) | 100
23 | SLOC 0.5 mg/L matrix spike | filter 1 |  |  |  | NPOC Desired Conc (mg/L) | 2
24 | SLOC plain matrix | filter 1
25 | SLOW 10 mg/L matrix spike | filter 4 |  |  |  | NPOC Vol Needed (mL) | 0.2 | 0.195
26 | SLOW 5 mg/L matrix spike | filter 4
27 | SLOW 2 mg/L matrix spike | filter 4 | rinsed with 50 mL plain SLOW
28 | SLOW 1 mg/L matrix spike | filter 3 |  |  |  | Stock Conc (ppm) | 1000
29 | SLOW 0.5 mg/L matrix spike | filter 3 |  |  |  | Vol Flask Final (mL) | 100
30 | SLOW plain matrix | filter 3 |  |  |  | NPOC Desired Conc (mg/L) | 1
31 | VDOS 10 mg/L matrix spike | filter 6
32 | VDOS 5 mg/L matrix spike | filter 6 |  |  |  | NPOC Vol Needed (mL) | 0.1 | 0.0956
33 | VDOS 2 mg/L matrix spike | filter 6 | rinsed with 50 mL plain VDOS
34 | VDOS 1 mg/L matrix spike | filter 5
35 | VDOS 0.5 mg/L matrix spike | filter 5 |  |  |  | Stock Conc (ppm) | 1000
36 | VDOS plain matrix | filter 5 |  |  |  | Vol Flask Final (mL) | 100
37 | VDOW 10 mg/L matrix spike | filter 8 |  |  |  | NPOC Desired Conc (mg/L) | 0.5
38 | VDOW 5 mg/L matrix spike | filter 8
39 | VDOW 2 mg/L matrix spike | filter 8 | rinsed with 50 mL plain VDOW |  |  | NPOC Vol Needed (mL) | 0.05 | 0.0484
40 | VDOW 1 mg/L matrix spike | filter 7
41 | VDOW 0.5 mg/L matrix spike | filter 7
42 | VDOW plain matrix | filter 7
43 | SLOC plain matrix | REP | filter 1 |  |  | Matrix Spike Concentration | 10
44 | SLOC 10 mg/L matrix spike | REP | filter 2 |  |  | Final Volume MS (mL) | 500
45 | SLOW plain matrix | REP | filter 3 |  |  | Conc. Spiking Solution | 1000
46 | SLOW 10 mg/L matrix spike | REP | filter 4 |  |  | Volume Spiking Solution | 5
47 | VDOS plain matrix | REP | filter 5
48 | VDOS 10 mg/L matrix spike | REP | filter 6 |  |  | Matrix Spike Concentration | 5
49 | VDOW plain matrix | REP | filter 7 |  |  | Final Volume MS (mL) | 250
50 | VDOW 10 mg/L matrix spike | REP | filter 8 |  |  | Conc. Spiking Solution | 1000
51 | filter test: sample 1 | 5 mL to wet + 5 mL pushed through |  |  |  | Volume Spiking Solution | 1.25
52 | filter test: sample 2 | 10-15 mL
53 | filter test: sample 3 | 15-20 mL |  |  |  | Matrix Spike Concentration | 2
54 | filter test: sample 4 | 20-25 mL |  |  |  | Final Volume MS (mL) | 250
55 | filter test: sample 5 | 25-30 mL |  |  |  | Conc. Spiking Solution | 1000
56 | filter test: sample 6 | 30-35 mL |  |  |  | Volume Spiking Solution | 0.5
57 | filter test: sample 7 | 35-40 mL
58 | filter test: sample 8 | 40-45 mL |  |  |  | Matrix Spike Concentration | 1
59 | filter test: sample 9 | 45-50 mL |  |  |  | Final Volume MS (mL) | 250
60 | filter test: sample 10 | refill syringe, 50-55 mL |  |  |  | Conc. Spiking Solution | 1000
61 | filter test: sample 11 | 55-60 mL |  |  |  | Volume Spiking Solution | 0.25
62 | filter test: sample 12 | 60-65 mL
63 | filter test: sample 13 | 65-70 mL |  |  |  | Matrix Spike Concentration | 0.5
64 | filter test: sample 14 | 70-75 mL |  |  |  | Final Volume MS (mL) | 500
65 | filter test: sample 15 | 75-80 mL |  |  |  | Conc. Spiking Solution | 1000
66 | filter test: sample 16 | 80-85 mL |  |  |  | Volume Spiking Solution | 0.25
67 | filter test: sample 17 | 85-90 mL
68 | filter test: sample 18 | 90-95 mL |  |  |  | % Recovery = [(spiked sample result - unspiked sample result) x 100% ] / known spike added concentration
69 | filter test: sample 19 | 95-100 mL
70 | filter test: sample 20 | refill syringe, 100-105 mL
71
72
73
74
75
76
77
78
79
80
81
82
83
84
85
86
87
88
89
90
91
92
93

## data_report
Sample_ID | NPOC_mg_L | TN_mg_L
SLOC 10 mg/L matrix spike | 12.8613
SLOC 5 mg/L matrix spike | 8.2439
SLOC 2 mg/L matrix spike | 5.5309
SLOC 1 mg/L matrix spike | 4.4144
SLOC 0.5 mg/L matrix spike | 4.0832
SLOC plain matrix | 3.4323
SLOW 10 mg/L matrix spike | 11.923
SLOW 5 mg/L matrix spike | 6.9436
SLOW 2 mg/L matrix spike | 4.1021
SLOW 1 mg/L matrix spike | 2.9872
SLOW 0.5 mg/L matrix spike | 2.4209
SLOW plain matrix | 1.9656
VDOS 10 mg/L matrix spike | 14.7687
VDOS 5 mg/L matrix spike | 9.7105
VDOS 2 mg/L matrix spike | 6.7816
VDOS 1 mg/L matrix spike | 5.8461
VDOS 0.5 mg/L matrix spike | 5.2915
VDOS plain matrix | 4.8858
VDOW 10 mg/L matrix spike | 13.9106
VDOW 5 mg/L matrix spike | 9.2742
VDOW 2 mg/L matrix spike | 6.443
VDOW 1 mg/L matrix spike | 5.5717
VDOW 0.5 mg/L matrix spike | 4.9865
VDOW plain matrix | 4.4436
filter test: sample 1 | 185.0054
filter test: sample 2 | 29.348
filter test: sample 3 | 20.4457
filter test: sample 4 | 8.4277
filter test: sample 5 | 4.7778
filter test: sample 6 | 2.653
filter test: sample 7 | 1.7905
filter test: sample 8 | 0.80175
filter test: sample 9 | 0.42261
filter test: sample 10 | 1.2634
filter test: sample 11 | 0.69507
filter test: sample 12 | 0.2882
filter test: sample 13 | 0.15992
filter test: sample 14 | 0.09232
filter test: sample 15 | 0.059995
filter test: sample 16 | 0.040147
filter test: sample 17 | 0.038134
filter test: sample 18 | 0.076953
filter test: sample 19 | 0.12183
filter test: sample 20 | 0.4264